# Temperature Datalogger Calibration Record

Internal calibration record for the single custom PT100 datalogger unit used in the test laboratory.

# Revision History

| Revision | Date | Description of Change | Approved By / Initials |
| --- | --- | --- | --- |
| 0 | 2026-03-06 | Initial editable calibration record issued for the single PT100 datalogger unit. | RS |
| 1 | 2026-03-06 | Separated assigned calibration uncertainty from in-use allowance and added basis notes for assigned values. | RS |

# Instrument Identification

| Instrument | PT100 Temperature Datalogger (custom single unit) |
| --- | --- |
| Controller | ESP32-S3 |
| RTD Interface | MAX31865 |
| Probe Model | Omega PR-21D-3-100-A-1/8-0300-M12-1 |
| Location | Bemidji Test Laboratory |
| Atmospheric Pressure Source | KBJI METAR (Bemidji Regional Airport, adjacent to site) |

# Measurement Configuration Snapshot

| Item | Value |
| --- | --- |
| RTD Type | PT100 |
| Wiring | 4-wire |
| Reference Resistor | 430 ohm |
| MAX31865 Filter | 60 Hz |
| MAX31865 Conversion Mode | PT100 table |
| Bias Settle Delay | 10 ms |
| Temperature Resolution | 0.03125 C nominal |
| Calibration Model | 2-point linear |
| Calibration Storage | ESP32 NVS |

# Calibration Procedure

Ice point procedure. Distilled water and crushed ice are prepared as a true slurry/slush, not floating ice in bulk water. The probe is immersed approximately 5 to 6 inches into the slush in a 2 L beaker. The probe is gently moved/stirred in the slurry during stabilization to minimize local pockets and maintain equilibrium at the sensing element. This point is assigned a reference temperature of 0.000 C.

Steam point procedure. Distilled water is boiled vigorously in a loosely covered beaker. The probe is placed several inches above the water surface within the steam space under the cover, without immersion in liquid water. The steam reference temperature is calculated in firmware from KBJI METAR pressure data by converting the reported altimeter setting to estimated station pressure and then calculating the saturation temperature of water at that pressure.

Model application. The raw measured temperatures at the two reference points are used to calculate and store a linear 2-point calibration model in nonvolatile memory within the datalogger.

# METAR Pressure Record

| Field | Record |
| --- | --- |
| METAR Station | KBJI - Bemidji Regional Airport |
| Calibration Date/Time | 2026-02-18T17:56:35Z (UTC) |
| METAR Observation Time | 11:35 AM |
| Elevation (FAA Chart Supplement) | 1391 feet |
| Altimeter Setting | 29.23 inHg |
| Estimated Station Pressure | 27.804 inHg |
| Calculated Steam Temperature | 97.967 C |
| Notes | Airport is adjacent to calibration site |

# Calibration Data and Stored Calibration Points

| Point | Raw Avg (C) | Reference (C) | Residual (C) |
| --- | --- | --- | --- |
| Ice Point | 0.471 | 0.000 | -0.471 |
| Steam Point | 98.206 | 97.967 | -0.239 |

## Coefficient Record

| Coefficient Record | Value |
| --- | --- |
| Point 1 Raw | 0.471 C |
| Point 1 Actual | 0.000 C |
| Point 2 Raw | 98.206 C |
| Point 2 Actual | 97.967 C |
| Calibration Mode | Linear |

# Assigned Calibration Uncertainty

Purpose. This section states the assigned calibration uncertainty of the instrument as calibrated. The values below are assigned internal working standard uncertainties based on manufacturer information, instrument design, the observed calibration method, and engineering judgment. They are not derived from a formal repeated-study uncertainty analysis.

| Calibration component | Std. unc. (C) | Brief basis |
| --- | --- | --- |
| Probe residual / interpolation allowance | 0.12 | Residual probe-model allowance after 2-point endpoint calibration. Based on the 1/3 DIN PT100 probe class, but smaller than the full uncalibrated end-of-range tolerance because endpoint offset and gain are already corrected. |
| Electronics / conversion / quantization | 0.03 | Residual MAX31865 readout contribution after calibration. Covers quantization plus small remaining reference/conversion effects; intentionally larger than bare LSB math alone. |
| Fixed-point realization (ice + steam) | 0.10 | Method allowance for the stirred slush ice bath and the pressure-corrected steam point. Dominated mainly by the practical steam setup and pressure-estimation method, not by the ice bath. |

Calculation. Combined calibration standard uncertainty uc,cal=0.122+ 0.032+ 0.102=0.159 °C, rounded to 0.16 C. Assigned calibration expanded uncertainty Ucal=k∙uc,cal with k=2=±0.32 °C, rounded to ±0.32 °C.

Meaning. When the logger is used as a calibrated instrument under conditions similar to the calibration setup, the indicated temperature is assigned a calibration uncertainty of approximately ±0.32 °C at k=2.

# Assigned In-Use Allowance and Working Uncertainty

Purpose. This section separates real-world use effects from the calibration itself so the record shows both the calibration-only result and the more conservative working uncertainty used in service.

| In-use component | Std. unc. (C) | Brief basis |
| --- | --- | --- |
| Placement / coupling / stem-conduction allowance | 0.08 | Allowance for probe placement, thermal coupling, local gradients, and stem-conduction effects during normal use. This term is not part of the fixed-point calibration itself; it is added to represent practical field/lab use. |

Calculation. Combined working standard uncertainty uc,work=0.1592+ 0.082=0.178 °C, rounded to 0.18 °C. Assigned working expanded uncertainty Uwork=k ∙ uc,work with k=2=±0.36 °C, rounded to ±0.36 °C.

Meaning. This is the more conservative value to use for actual decisions, comparisons, and audit discussions when the instrument is being used outside the calibration setup.

# Traceability and Use Statement

This calibration is an internally controlled working calibration based on physical temperature reference points. The atmospheric pressure input for the steam point is derived from the adjacent airport METAR record and converted in firmware to estimated station pressure. This record supports internal verification and audit documentation for the single custom datalogger unit. It is not an accredited laboratory calibration certificate.

# Authorization

| Technician | Robert Susmilch |
| --- | --- |
| Calibration Date | February 18, 2026 |
| Next Calibration Due | 1 year from calibration date |