# Temperature Datalogger Calibration Record

Internal calibration record for the PT100 datalogger.

# Revision History

| Revision | Date | Description of Change | Approved By / Initials |
| --- | --- | --- | --- |
| 0 | 2026-03-06 | Initial editable calibration record issued for the single PT100 datalogger unit. | RS |
| 1 | 2026-03-06 | Separated assigned calibration uncertainty from in-use allowance and added basis notes for assigned values. | RS |
| 2 | 2026-04-08 | Updated template for Omega PR-21D-3-100-A-1/8-0300-M12-1, resistance-domain fixed-point calibration workflow, steam-column method, and revised uncertainty structure. | RS |

# Instrument Identification

| Item | Value |
| --- | --- |
| Instrument | PT100 Temperature Datalogger |
| Calibration ID |  |
| Internal ID (MAC address) | 98:A3:16:12:57:B0 |
| Controller | ESP32-S3 |
| RTD Interface | MAX31865 |
| Probe Model | Omega PR-21D-3-100-A-1/8-0300-M12-1 |
| Probe Construction Notes | Pt100, Class A, 4-wire, 316L sheath, 1/8 in diameter, 3 in length, M12 connection |
| Location | Bemidji Test Laboratory |
| Atmospheric Pressure Source | KBJI METAR (Bemidji Regional Airport, adjacent to site) or local calibrated barometer if available |
| Calibration Domain in Firmware | Resistance-domain fixed-point model preferred (raw measured ohms -> corrected ohms -> corrected temperature) |

# Measurement Configuration Snapshot

| Item | Value |
| --- | --- |
| RTD Type | PT100 |
| Wiring | 4-wire |
| Reference Resistor | 430 ohm |
| MAX31865 Filter | 60 Hz |
| MAX31865 Conversion Mode | PT100 table |
| Bias Settle Delay | 10 ms |
| Temperature Resolution | 0.03125 C nominal |
| Calibration Model | 2-point linear unless engineering change approves otherwise |
| Calibration Storage | ESP32 NVS |
| Record Note | Template should retain both temperature-domain values for readability and resistance-domain values for audit traceability. |

# Calibration Procedure

Ice point procedure. Prepare distilled water and crushed ice as a true slush, not floating ice in bulk water. Immerse the sensing section deeply enough that the active element is fully surrounded by slurry. Gently move or stir the probe during stabilization to break up local pockets. Capture the point at 0.000 C.

Steam point procedure. Boil distilled water vigorously in a vented beaker and place the probe in the attached steam column so the sensing section is fully inside the saturated steam path and not in liquid water. Keep the column vented; do not seal the system. The steam reference temperature is calculated in firmware using satpt from either local station pressure or METAR altimeter plus site elevation. Record the exact pressure input used.

Model application. The firmware prefers resistance-domain calibration. During capture, the instrument stores raw average temperature, raw average resistance, and their stability statistics. During calibration application, the entered reference temperature is converted to ideal Pt100 resistance and the fit is performed in ohms; runtime correction then applies in ohms before converting back to temperature.

# Pressure / Steam Reference Record

| Field | Record |
| --- | --- |
| Pressure source type | METAR |
| METAR Station / Barometer ID | KBJI - Bemidji Regional Airport |
| Calibration Date/Time (UTC) | 2026-04-08T20:12:33Z |
| Observation Time | 2026-04-08T19:55:00Z |
| Elevation used | 1391 feet |
| Altimeter Setting | A2969 |
| Estimated Station Pressure | 28.23 inHg |
| Calculated Steam Temperature | 98.379 |
| Notes | Probe centered in steam path; vent remained open. |

# Calibration Data and Stored Calibration Points

Use the values actually reported by the firmware window/capture data. The record should preserve both operator-readable temperature residuals and the resistance-domain values actually used for fitting.

| Point | Raw Avg (C) | Raw Avg (Ohm) | Reference (C) | Ideal Ref (Ohm) | Residual (C) | Residual (Ohm) | Std. Dev. (C) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Ice Point | [ ] | [ ] | 0.000 | [ ] | [Reference - Raw Avg C] | [Ideal - Raw Avg Ohm] | [ ] |
| Steam Point | [ ] | [ ] | [satpt result] | [ ] | [Reference - Raw Avg C] | [Ideal - Raw Avg Ohm] | [ ] |

# Stored Model / Coefficient Record

| Coefficient Record | Value |
| --- | --- |
| Calibration mode | [linear / piecewise / polyN] |
| Calibration domain | Resistance |
| Point 1 raw temperature | [ ] C |
| Point 1 raw resistance | [ ] ohm |
| Point 1 reference temperature | 0.000 C |
| Point 1 ideal reference resistance | [ ] ohm |
| Point 2 raw temperature | [ ] C |
| Point 2 raw resistance | [ ] ohm |
| Point 2 reference temperature / ideal resistance | [ ] C / [ ] ohm |

# Assigned Calibration Uncertainty

Purpose. This section states the assigned calibration uncertainty of the instrument as calibrated. The values below are recommended working values for this updated hardware and steam-column method. They should be confirmed or adjusted after the first few completed calibrations and after reviewing actual residuals.

| Calibration component | Std. unc. (C) | Brief basis |
| --- | --- | --- |
| Probe residual / interpolation allowance | 0.08 | Recommended starting value for the Class A probe after 2-point endpoint calibration in resistance domain. Tightened from the earlier template because the probe class is better and the fit is applied in ohms, but still leaves room for interpolation and element/sheath behavior. |
| Electronics / conversion / quantization | 0.03 | Carryover value for residual MAX31865 readout, reference resistor, conversion, and quantization effects after calibration. |
| Fixed-point realization (ice + steam column) | 0.07 | Reduced from the earlier steam-space allowance because the attached steam column gives a better defined saturated-vapor zone and reduces splash / surface-gradient ambiguity, while still keeping allowance for pressure estimation, venting, equilibration, and practical setup scatter. |

Calculation. Combined calibration standard uncertainty u_c = sqrt(0.08^2 + 0.03^2 + 0.07^2) = 0.110 C, rounded to 0.11 C. Assigned calibration expanded uncertainty U = k x u_c with k = 2, rounded to 0.22 C.

Meaning. When the logger is used as a calibrated instrument under conditions similar to the calibration setup, the indicated temperature is provisionally assigned a calibration uncertainty of approximately +/-0.22 C at k = 2. Replace this text if a different uncertainty budget is justified by actual study data.

# Assigned In-Use Allowance and Working Uncertainty

Purpose. This section separates real-world use effects from the calibration itself so the record shows both the calibration-only result and the more conservative working uncertainty used in service.

| In-use component | Std. unc. (C) | Brief basis |
| --- | --- | --- |
| Placement / coupling / stem-conduction allowance | 0.08 | Retained as a practical allowance for probe placement, thermal coupling, local gradients, insertion depth sensitivity, and stem-conduction effects during normal use. The shorter 3 in sheath does not justify relaxing this without actual use data. |

Calculation. Combined working standard uncertainty u_w = sqrt(0.110^2 + 0.08^2) = 0.136 C, rounded to 0.14 C. Assigned working expanded uncertainty U = k x u_w with k = 2, rounded to 0.27 C.

Meaning. This is the more conservative value to use for actual decisions, comparisons, and audit discussions when the instrument is being used outside the calibration setup.

# Residual Review Notes

Residuals should now be reviewed in both Celsius and ohms. The Celsius residual remains useful for the record, but the ohmic residual better matches the actual fit domain used by the firmware.

If a calibrated local barometer is available at the bench, it is preferable to METAR-derived pressure for the steam point and may justify a smaller steam-method allowance than 0.07 C.

# Traceability and Use Statement

This calibration is an internally controlled working calibration based on physical temperature reference points. The steam point reference may be derived either from a local barometer or from adjacent-airport METAR data converted in firmware to estimated station pressure. This record supports internal verification and audit documentation for the single custom datalogger unit. It is not an accredited laboratory calibration certificate.

# Authorization

| Field | Value |
| --- | --- |
| Technician | Robert Susmilch |
| Calibration Date | [Insert date] |
| Next Calibration Due | 1 year from calibration date unless shortened by policy or observed drift |